# Dataset metrics and scanner comparisons

Table 2. Paired baseline-to-clean changes by scanner and specimen cohort.

| Cohort | Scanner | ΔE₀₀ | ΔRMS height 100 µm grid (µm) | ΔMedian native slope (°) | Triangle density clean/baseline |
| --- | --- | --- | --- | --- | --- |
| Fresh n = 16 | iTERO | 8.70 [7.10, 9.95] | 1.51 [−0.39, 7.07] | −0.15 [−0.31, −0.07] | 0.73 [0.65, 0.86] |
| Fresh n = 16 | TRIOS3 | 6.12 [5.76, 6.34] | −0.23 [−1.14, 0.49] | −0.29 [−0.36, −0.20] | 0.65 [0.59, 0.73] |
| Fresh n = 16 | TRIOS5 | 7.66 [6.20, 8.72] | −0.50 [−1.58, 0.58] | −0.31 [−0.44, −0.17] | 0.62 [0.56, 0.71] |
| Fresh n = 16 | LABscanner | 2.94 [2.48, 3.47] | −1.35 [−2.43, −0.85] | −2.29 [−2.47, −1.98] | 0.55 [0.53, 0.59] |
| Mature n = 32 | iTERO | 8.24 [7.26, 9.80] | −0.28 [−1.87, 1.06] | −0.36 [−0.50, −0.29] | 0.66 [0.58, 0.70] |
| Mature n = 32 | TRIOS3 | 6.88 [6.27, 7.37] | −0.66 [−1.03, −0.20] | −0.29 [−0.35, −0.25] | 0.69 [0.66, 0.71] |
| Mature n = 32 | TRIOS5 | 9.34 [8.60, 9.88] | −0.81 [−1.23, −0.16] | −0.49 [−0.58, −0.38] | 0.57 [0.54, 0.62] |
| Mature n = 32 | LABscanner | 2.93 [2.51, 3.18] | −1.43 [−2.01, −0.92] | −1.98 [−2.08, −1.72] | 0.55 [0.52, 0.59] |

Values are medians with 95% bootstrap confidence intervals in brackets. Changes are clean minus baseline; triangle density is the clean/baseline ratio. Negative height or slope changes indicate a reduction. Each comparison contains 16 fresh or 32 mature independent specimens. Intervals use 1,000 specimen-level bootstrap resamples. ΔE₀₀ compares area-weighted scan-mean CIELAB colours. RMS height is the root-mean-square plane-detrended height after averaging into 100 µm cells.

Table S2 provides the descriptive values underlying the comparisons, including metrics not displayed in Figures 4–7. In the manuscript, cite Table 2 for paired stage changes and Supplementary Table S2 for colour-channel values, mesh counts, edge lengths and height summaries.

# Table S2A Fresh specimens iTERO

Values are medians [P25–P75] across 16 specimens per stage. Counts and triangle areas are rounded to whole units.

| Metric | Baseline | Clean |
| --- | --- | --- |
| Mean L* | 47.67 [45.38–48.40] | 55.34 [54.15–55.88] |
| Within-scan SD L* | 3.35 [2.94–3.85] | 2.28 [2.12–2.75] |
| Median L* | 47.54 [46.50–49.24] | 55.56 [54.75–56.10] |
| P05 L* | 41.58 [39.15–43.40] | 50.19 [49.16–52.19] |
| P95 L* | 51.80 [50.44–53.68] | 58.63 [57.77–59.42] |
| Mean a* | 50.33 [48.86–51.71] | 39.99 [38.68–40.74] |
| Within-scan SD a* | 2.66 [1.98–3.08] | 2.61 [2.05–2.98] |
| Median a* | 50.66 [48.80–51.74] | 40.65 [39.26–41.51] |
| P05 a* | 45.78 [44.25–47.41] | 35.19 [33.08–35.93] |
| P95 a* | 53.81 [52.64–55.29] | 43.07 [42.20–44.07] |
| Mean b* | 17.60 [15.95–18.73] | 12.49 [11.91–12.99] |
| Within-scan SD b* | 2.62 [2.39–2.99] | 1.38 [1.25–1.97] |
| Median b* | 17.78 [16.02–18.74] | 12.72 [11.84–13.10] |
| P05 b* | 13.83 [11.85–14.74] | 9.50 [8.99–10.75] |
| P95 b* | 21.06 [19.55–23.68] | 15.00 [14.15–15.27] |
| Encoded vertices | 2566 [2419–2692] | 1854 [1627–2108] |
| Unique geometric vertices | 2086 [2030–2220] | 1542 [1438–1782] |
| Triangle count | 3933 [3839–4190] | 2862 [2684–3322] |
| Mesh surface area (mm²) | 111.09 [110.99–111.27] | 110.43 [110.15–110.70] |
| Encoded vertex density (mm⁻²) | 23.07 [21.75–24.21] | 16.81 [14.74–19.09] |
| Geometric vertex density (mm⁻²) | 18.73 [18.27–19.99] | 14.18 [13.06–16.08] |
| Triangle density (mm⁻²) | 35.31 [34.58–37.70] | 26.31 [24.35–29.97] |
| Triangle area P10 (µm²) | 11026 [10221–12094] | 15140 [12560–17547] |
| Triangle area median (µm²) | 25247 [23618–25866] | 34583 [29815–37169] |
| Triangle area P90 (µm²) | 48457 [45643–51442] | 65052 [58693–71921] |
| Edge length P10 (µm) | 147.9 [144.2–154.0] | 170.8 [155.9–182.4] |
| Edge length median (µm) | 252.5 [245.9–257.9] | 295.5 [279.8–310.8] |
| Edge length P90 (µm) | 394.2 [385.8–404.1] | 461.7 [435.2–484.1] |
| Native mean absolute height (µm) | 4.45 [4.07–5.34] | 6.09 [4.20–10.17] |
| Native RMS height (µm) | 5.89 [5.60–7.32] | 8.46 [5.80–13.47] |
| Native height P95–P5 (µm) | 18.42 [16.83–22.90] | 25.13 [17.35–38.92] |
| Native face slope median (°) | 0.98 [0.88–1.01] | 0.77 [0.70–0.89] |
| Native face slope P95 (°) | 2.88 [2.75–3.22] | 2.45 [1.81–3.68] |
| Mean absolute height, 50 µm grid (µm) | 4.41 [4.00–5.29] | 5.99 [4.12–10.08] |
| RMS height, 50 µm grid (µm) | 5.81 [5.42–7.16] | 8.18 [5.53–13.23] |
| Height P95–P5, 50 µm grid (µm) | 18.26 [16.57–22.68] | 24.89 [17.31–38.56] |
| Median grid slope, 50 µm grid (°) | 0.85 [0.76–0.87] | 0.68 [0.60–0.81] |
| Mean absolute height, 100 µm grid (µm) | 4.36 [3.93–5.24] | 5.93 [4.06–9.97] |
| RMS height, 100 µm grid (µm) | 5.71 [5.31–7.07] | 8.02 [5.28–12.99] |
| Height P95–P5, 100 µm grid (µm) | 18.03 [16.30–22.39] | 24.59 [17.25–38.23] |
| Median grid slope, 100 µm grid (°) | 0.74 [0.66–0.76] | 0.59 [0.53–0.73] |
| Mean absolute height, 200 µm grid (µm) | 4.25 [3.77–5.12] | 5.82 [3.98–9.74] |
| RMS height, 200 µm grid (µm) | 5.50 [5.06–6.89] | 7.77 [5.09–12.55] |
| Height P95–P5, 200 µm grid (µm) | 17.65 [15.74–21.86] | 24.07 [17.04–37.50] |
| Median grid slope, 200 µm grid (°) | 0.55 [0.50–0.58] | 0.48 [0.40–0.61] |

P05/P95 and SD refer to within-scan colour distributions; native edge and triangle percentiles are also calculated within scans. Table cells then summarise those scan-level values across specimens. Grid descriptors use common physical cell averaging; they do not equalise scanner resolution. All heights are plane-detrended; colours are uncalibrated.

# Table S2B Fresh specimens TRIOS3

Values are medians [P25–P75] across 16 specimens per stage. Counts and triangle areas are rounded to whole units.

| Metric | Baseline | Clean |
| --- | --- | --- |
| Mean L* | 49.62 [49.28–50.08] | 54.45 [54.24–54.97] |
| Within-scan SD L* | 1.92 [1.71–2.25] | 0.64 [0.59–0.71] |
| Median L* | 49.95 [49.21–50.15] | 54.49 [54.31–54.93] |
| P05 L* | 46.36 [45.79–46.64] | 53.38 [52.85–53.99] |
| P95 L* | 52.53 [52.17–52.82] | 55.41 [55.06–56.00] |
| Mean a* | 45.13 [44.70–46.11] | 35.63 [35.37–36.44] |
| Within-scan SD a* | 1.25 [1.12–1.44] | 1.48 [1.07–1.60] |
| Median a* | 44.96 [44.71–45.94] | 35.72 [35.58–36.07] |
| P05 a* | 43.34 [42.81–43.96] | 33.61 [31.91–34.55] |
| P95 a* | 47.30 [46.65–48.25] | 37.36 [36.62–38.85] |
| Mean b* | 23.03 [21.84–24.47] | 18.44 [18.24–18.69] |
| Within-scan SD b* | 2.55 [2.18–2.87] | 1.24 [1.14–1.53] |
| Median b* | 23.12 [21.88–24.52] | 18.99 [18.65–19.14] |
| P05 b* | 18.92 [18.00–20.51] | 16.13 [15.22–16.67] |
| P95 b* | 26.61 [25.37–29.31] | 19.76 [19.39–20.43] |
| Encoded vertices | 598 [554–632] | 402 [371–441] |
| Unique geometric vertices | 598 [554–632] | 402 [371–441] |
| Triangle count | 1079 [1004–1151] | 702 [643–772] |
| Mesh surface area (mm²) | 110.93 [110.31–111.34] | 111.04 [110.27–111.40] |
| Encoded vertex density (mm⁻²) | 5.40 [5.02–5.68] | 3.63 [3.33–3.96] |
| Geometric vertex density (mm⁻²) | 5.40 [5.02–5.68] | 3.63 [3.33–3.96] |
| Triangle density (mm⁻²) | 9.71 [9.10–10.36] | 6.34 [5.81–6.93] |
| Triangle area P10 (µm²) | 24739 [21040–31586] | 17480 [13756–21735] |
| Triangle area median (µm²) | 99717 [92314–108459] | 154226 [141121–170721] |
| Triangle area P90 (µm²) | 179697 [161445–189878] | 306014 [286177–314352] |
| Edge length P10 (µm) | 278.6 [254.8–295.4] | 249.6 [207.1–278.7] |
| Edge length median (µm) | 489.3 [480.9–516.7] | 603.1 [578.7–637.2] |
| Edge length P90 (µm) | 738.9 [692.7–756.5] | 965.4 [930.4–976.6] |
| Native mean absolute height (µm) | 3.53 [3.01–4.80] | 3.07 [2.56–5.20] |
| Native RMS height (µm) | 4.55 [3.92–6.16] | 3.91 [3.21–6.56] |
| Native height P95–P5 (µm) | 14.79 [13.07–19.90] | 12.85 [10.46–20.71] |
| Native face slope median (°) | 0.48 [0.44–0.59] | 0.23 [0.19–0.26] |
| Native face slope P95 (°) | 1.65 [1.55–2.16] | 0.66 [0.57–0.84] |
| Mean absolute height, 50 µm grid (µm) | 3.52 [2.99–4.77] | 3.06 [2.55–5.16] |
| RMS height, 50 µm grid (µm) | 4.51 [3.89–6.11] | 3.87 [3.19–6.51] |
| Height P95–P5, 50 µm grid (µm) | 14.73 [13.02–19.86] | 12.76 [10.43–20.69] |
| Median grid slope, 50 µm grid (°) | 0.44 [0.41–0.54] | 0.22 [0.18–0.25] |
| Mean absolute height, 100 µm grid (µm) | 3.50 [2.97–4.73] | 3.05 [2.54–5.12] |
| RMS height, 100 µm grid (µm) | 4.47 [3.86–6.06] | 3.83 [3.15–6.45] |
| Height P95–P5, 100 µm grid (µm) | 14.66 [12.90–19.74] | 12.72 [10.39–20.61] |
| Median grid slope, 100 µm grid (°) | 0.40 [0.37–0.47] | 0.21 [0.18–0.24] |
| Mean absolute height, 200 µm grid (µm) | 3.44 [2.91–4.63] | 3.03 [2.52–5.06] |
| RMS height, 200 µm grid (µm) | 4.36 [3.77–5.93] | 3.79 [3.11–6.34] |
| Height P95–P5, 200 µm grid (µm) | 14.38 [12.53–19.45] | 12.62 [10.28–20.44] |
| Median grid slope, 200 µm grid (°) | 0.33 [0.30–0.37] | 0.19 [0.16–0.23] |

P05/P95 and SD refer to within-scan colour distributions; native edge and triangle percentiles are also calculated within scans. Table cells then summarise those scan-level values across specimens. Grid descriptors use common physical cell averaging; they do not equalise scanner resolution. All heights are plane-detrended; colours are uncalibrated.

# Table S2C Fresh specimens TRIOS5

Values are medians [P25–P75] across 16 specimens per stage. Counts and triangle areas are rounded to whole units.

| Metric | Baseline | Clean |
| --- | --- | --- |
| Mean L* | 51.23 [50.25–52.57] | 57.13 [56.95–57.96] |
| Within-scan SD L* | 2.94 [2.57–3.11] | 0.84 [0.72–0.94] |
| Median L* | 51.63 [50.50–52.95] | 57.28 [57.09–57.93] |
| P05 L* | 46.40 [45.61–47.12] | 55.59 [55.19–56.45] |
| P95 L* | 55.75 [54.80–57.67] | 58.35 [58.09–58.99] |
| Mean a* | 38.22 [37.01–39.53] | 27.41 [26.64–28.01] |
| Within-scan SD a* | 2.23 [2.06–2.69] | 1.57 [1.36–1.71] |
| Median a* | 38.37 [37.40–39.71] | 27.44 [26.85–27.94] |
| P05 a* | 33.91 [31.58–35.60] | 24.62 [23.71–25.39] |
| P95 a* | 41.92 [40.98–42.61] | 29.40 [28.90–29.95] |
| Mean b* | 17.18 [16.05–17.89] | 12.25 [11.42–12.56] |
| Within-scan SD b* | 2.04 [1.62–2.52] | 1.28 [1.06–1.40] |
| Median b* | 17.20 [16.03–17.80] | 12.71 [11.24–12.84] |
| P05 b* | 13.74 [13.05–15.06] | 9.87 [9.71–10.40] |
| P95 b* | 20.16 [19.41–21.39] | 14.01 [13.50–14.40] |
| Encoded vertices | 552 [504–594] | 336 [323–368] |
| Unique geometric vertices | 552 [504–594] | 336 [323–368] |
| Triangle count | 1002 [912–1081] | 588 [568–653] |
| Mesh surface area (mm²) | 104.20 [103.39–104.93] | 104.54 [103.65–105.37] |
| Encoded vertex density (mm⁻²) | 5.30 [4.92–5.58] | 3.24 [3.08–3.54] |
| Geometric vertex density (mm⁻²) | 5.30 [4.92–5.58] | 3.24 [3.08–3.54] |
| Triangle density (mm⁻²) | 9.62 [8.90–10.17] | 5.66 [5.45–6.27] |
| Triangle area P10 (µm²) | 49854 [41268–59917] | 74502 [61253–85937] |
| Triangle area median (µm²) | 97390 [91465–106829] | 167186 [153813–175176] |
| Triangle area P90 (µm²) | 164748 [157050–181196] | 281753 [255097–297351] |
| Edge length P10 (µm) | 328.9 [298.1–353.4] | 406.5 [369.1–421.2] |
| Edge length median (µm) | 495.1 [479.2–513.5] | 647.8 [611.3–659.1] |
| Edge length P90 (µm) | 715.6 [698.0–746.4] | 929.6 [887.9–965.3] |
| Native mean absolute height (µm) | 3.91 [3.31–4.98] | 2.99 [2.50–4.44] |
| Native RMS height (µm) | 5.05 [4.16–6.20] | 3.88 [3.08–5.63] |
| Native height P95–P5 (µm) | 16.26 [13.18–19.87] | 13.39 [10.31–19.05] |
| Native face slope median (°) | 0.57 [0.51–0.75] | 0.28 [0.23–0.30] |
| Native face slope P95 (°) | 2.10 [1.53–3.04] | 0.81 [0.65–0.85] |
| Mean absolute height, 50 µm grid (µm) | 3.90 [3.30–4.90] | 2.99 [2.50–4.39] |
| RMS height, 50 µm grid (µm) | 5.02 [4.14–6.16] | 3.87 [3.07–5.60] |
| Height P95–P5, 50 µm grid (µm) | 16.14 [13.07–19.59] | 13.32 [10.29–18.84] |
| Median grid slope, 50 µm grid (°) | 0.51 [0.45–0.67] | 0.27 [0.22–0.29] |
| Mean absolute height, 100 µm grid (µm) | 3.86 [3.28–4.86] | 2.98 [2.49–4.36] |
| RMS height, 100 µm grid (µm) | 4.95 [4.10–6.12] | 3.85 [3.06–5.58] |
| Height P95–P5, 100 µm grid (µm) | 16.03 [12.95–19.47] | 13.27 [10.25–18.78] |
| Median grid slope, 100 µm grid (°) | 0.46 [0.41–0.60] | 0.26 [0.21–0.28] |
| Mean absolute height, 200 µm grid (µm) | 3.76 [3.23–4.76] | 2.96 [2.47–4.30] |
| RMS height, 200 µm grid (µm) | 4.79 [4.01–6.02] | 3.80 [3.05–5.50] |
| Height P95–P5, 200 µm grid (µm) | 15.69 [12.80–19.39] | 13.15 [10.16–18.37] |
| Median grid slope, 200 µm grid (°) | 0.37 [0.33–0.47] | 0.23 [0.20–0.26] |

P05/P95 and SD refer to within-scan colour distributions; native edge and triangle percentiles are also calculated within scans. Table cells then summarise those scan-level values across specimens. Grid descriptors use common physical cell averaging; they do not equalise scanner resolution. All heights are plane-detrended; colours are uncalibrated.

# Table S2D Fresh specimens LABscanner

Values are medians [P25–P75] across 16 specimens per stage. Counts and triangle areas are rounded to whole units.

| Metric | Baseline | Clean |
| --- | --- | --- |
| Mean L* | 55.33 [54.52–56.55] | 55.84 [54.59–56.55] |
| Within-scan SD L* | 1.83 [1.32–2.23] | 1.82 [1.67–2.09] |
| Median L* | 55.54 [54.25–56.51] | 55.82 [54.50–56.41] |
| P05 L* | 51.88 [51.31–52.92] | 52.85 [52.20–53.46] |
| P95 L* | 58.34 [56.98–59.70] | 58.77 [58.03–59.71] |
| Mean a* | 34.97 [33.91–35.97] | 29.79 [28.58–30.34] |
| Within-scan SD a* | 2.84 [2.58–3.10] | 1.94 [1.78–2.09] |
| Median a* | 35.55 [34.39–36.63] | 29.97 [28.78–30.79] |
| P05 a* | 29.50 [28.85–30.49] | 25.38 [24.93–26.77] |
| P95 a* | 39.05 [37.76–39.40] | 32.41 [31.06–32.74] |
| Mean b* | 27.34 [25.92–28.36] | 25.24 [24.54–26.44] |
| Within-scan SD b* | 2.87 [2.42–3.36] | 2.60 [2.40–2.75] |
| Median b* | 27.63 [25.96–28.47] | 26.00 [24.63–26.32] |
| P05 b* | 21.44 [20.40–22.94] | 20.26 [20.03–21.19] |
| P95 b* | 31.24 [29.04–32.11] | 29.39 [28.47–29.65] |
| Encoded vertices | 8454 [7757–9089] | 4569 [4436–4755] |
| Unique geometric vertices | 8454 [7757–9089] | 4569 [4436–4755] |
| Triangle count | 16552 [15154–17810] | 8870 [8614–9234] |
| Mesh surface area (mm²) | 106.18 [106.04–106.28] | 104.96 [104.82–105.09] |
| Encoded vertex density (mm⁻²) | 79.64 [73.17–85.53] | 43.53 [42.18–45.33] |
| Geometric vertex density (mm⁻²) | 79.64 [73.17–85.53] | 43.53 [42.18–45.33] |
| Triangle density (mm⁻²) | 155.91 [142.94–167.59] | 84.50 [81.91–88.03] |
| Triangle area P10 (µm²) | 3227 [3210–3273] | 6099 [5863–6357] |
| Triangle area median (µm²) | 5696 [5371–6194] | 11023 [10534–11277] |
| Triangle area P90 (µm²) | 10838 [9837–11827] | 18620 [17855–19169] |
| Edge length P10 (µm) | 80.3 [80.1–80.9] | 112.3 [108.7–114.3] |
| Edge length median (µm) | 121.8 [116.1–130.3] | 171.4 [168.6–173.9] |
| Edge length P90 (µm) | 191.6 [183.6–200.1] | 251.1 [247.7–256.0] |
| Native mean absolute height (µm) | 5.61 [5.19–5.96] | 4.09 [3.53–5.14] |
| Native RMS height (µm) | 7.31 [6.63–7.68] | 5.23 [4.48–6.40] |
| Native height P95–P5 (µm) | 23.61 [21.66–24.49] | 17.19 [14.59–21.27] |
| Native face slope median (°) | 4.01 [3.70–4.22] | 1.71 [1.60–1.80] |
| Native face slope P95 (°) | 10.78 [10.15–11.65] | 3.71 [3.49–3.80] |
| Mean absolute height, 50 µm grid (µm) | 5.49 [5.04–5.85] | 4.07 [3.49–5.10] |
| RMS height, 50 µm grid (µm) | 7.12 [6.47–7.53] | 5.18 [4.45–6.34] |
| Height P95–P5, 50 µm grid (µm) | 23.19 [21.04–24.14] | 17.10 [14.41–21.00] |
| Median grid slope, 50 µm grid (°) | 3.06 [2.83–3.22] | 1.31 [1.23–1.36] |
| Mean absolute height, 100 µm grid (µm) | 5.22 [4.74–5.65] | 4.00 [3.43–5.02] |
| RMS height, 100 µm grid (µm) | 6.71 [6.10–7.23] | 5.10 [4.39–6.26] |
| Height P95–P5, 100 µm grid (µm) | 21.90 [19.75–23.14] | 16.93 [14.23–20.74] |
| Median grid slope, 100 µm grid (°) | 2.10 [1.94–2.21] | 0.96 [0.90–0.98] |
| Mean absolute height, 200 µm grid (µm) | 4.40 [3.97–5.15] | 3.82 [3.28–4.87] |
| RMS height, 200 µm grid (µm) | 5.73 [5.10–6.46] | 4.87 [4.22–6.08] |
| Height P95–P5, 200 µm grid (µm) | 18.62 [16.74–20.79] | 16.33 [13.70–19.86] |
| Median grid slope, 200 µm grid (°) | 0.86 [0.83–0.93] | 0.51 [0.48–0.56] |

P05/P95 and SD refer to within-scan colour distributions; native edge and triangle percentiles are also calculated within scans. Table cells then summarise those scan-level values across specimens. Grid descriptors use common physical cell averaging; they do not equalise scanner resolution. All heights are plane-detrended; colours are uncalibrated.

# Table S2E Mature specimens iTERO

Values are medians [P25–P75] across 32 specimens per stage. Counts and triangle areas are rounded to whole units.

| Metric | Baseline | Partial | Clean |
| --- | --- | --- | --- |
| Mean L* | 47.83 [46.63–49.53] | 52.48 [50.67–54.16] | 55.33 [53.51–56.88] |
| Within-scan SD L* | 3.08 [2.62–3.61] | 2.79 [2.43–3.54] | 2.07 [1.67–2.64] |
| Median L* | 48.23 [46.59–49.87] | 53.24 [51.35–54.49] | 55.62 [53.84–57.06] |
| P05 L* | 42.43 [41.18–43.96] | 47.52 [44.58–49.31] | 51.34 [49.84–52.56] |
| P95 L* | 52.82 [51.57–54.22] | 56.35 [54.76–57.33] | 58.10 [57.35–60.06] |
| Mean a* | 56.25 [53.39–57.67] | 46.29 [44.42–48.18] | 41.71 [40.52–42.75] |
| Within-scan SD a* | 2.42 [2.02–2.98] | 4.25 [3.91–4.79] | 2.67 [2.28–3.09] |
| Median a* | 56.43 [53.77–57.69] | 46.31 [44.24–48.68] | 42.39 [41.26–43.31] |
| P05 a* | 51.42 [49.05–53.05] | 39.08 [37.60–41.35] | 36.62 [35.24–38.24] |
| P95 a* | 59.85 [56.88–61.45] | 54.52 [51.23–55.89] | 45.14 [43.98–46.56] |
| Mean b* | 13.56 [12.11–14.90] | 11.91 [11.37–12.48] | 11.99 [11.15–12.63] |
| Within-scan SD b* | 2.19 [1.89–2.43] | 1.87 [1.61–2.23] | 1.26 [1.04–1.50] |
| Median b* | 13.50 [12.14–14.97] | 11.91 [11.46–12.45] | 12.11 [11.26–12.83] |
| P05 b* | 9.75 [8.95–10.87] | 8.86 [8.41–9.32] | 9.51 [8.88–10.32] |
| P95 b* | 17.38 [15.81–17.85] | 14.70 [13.63–15.53] | 13.62 [12.83–14.51] |
| Encoded vertices | 2691 [2566–3008] | 1828 [1602–2117] | 1682 [1530–2055] |
| Unique geometric vertices | 2398 [2254–2746] | 1540 [1423–1826] | 1560 [1352–1762] |
| Triangle count | 4548 [4291–5266] | 2859 [2642–3422] | 2904 [2505–3280] |
| Mesh surface area (mm²) | 111.31 [110.92–111.56] | 110.74 [110.12–111.14] | 110.63 [110.31–111.11] |
| Encoded vertex density (mm⁻²) | 24.13 [22.95–27.01] | 16.50 [14.52–18.97] | 15.20 [13.77–18.54] |
| Geometric vertex density (mm⁻²) | 21.38 [20.35–24.73] | 13.91 [12.91–16.33] | 14.18 [12.27–15.85] |
| Triangle density (mm⁻²) | 40.70 [38.83–47.41] | 25.85 [23.93–30.60] | 26.44 [22.76–29.50] |
| Triangle area P10 (µm²) | 11017 [8389–11739] | 16239 [12043–18046] | 14357 [11998–17609] |
| Triangle area median (µm²) | 22285 [19623–23585] | 34388 [29522–38473] | 33436 [30653–39285] |
| Triangle area P90 (µm²) | 40411 [34986–43854] | 65247 [57485–70308] | 66104 [57222–77195] |
| Edge length P10 (µm) | 145.0 [132.4–152.0] | 175.9 [155.0–185.7] | 170.1 [157.3–182.2] |
| Edge length median (µm) | 239.6 [225.4–246.3] | 297.3 [272.7–316.2] | 292.5 [281.5–317.5] |
| Edge length P90 (µm) | 363.2 [336.7–375.6] | 462.3 [430.8–476.7] | 465.5 [434.6–499.4] |
| Native mean absolute height (µm) | 5.63 [4.40–8.49] | 4.83 [3.88–8.37] | 5.74 [4.23–8.08] |
| Native RMS height (µm) | 7.56 [5.84–11.23] | 6.76 [5.57–12.51] | 7.63 [6.25–11.62] |
| Native height P95–P5 (µm) | 23.17 [18.15–35.53] | 18.62 [15.80–35.41] | 22.35 [17.30–35.90] |
| Native face slope median (°) | 1.11 [1.01–1.33] | 0.76 [0.71–0.96] | 0.76 [0.68–0.82] |
| Native face slope P95 (°) | 3.38 [2.75–4.29] | 1.97 [1.82–3.36] | 2.21 [1.86–3.08] |
| Mean absolute height, 50 µm grid (µm) | 5.57 [4.25–8.42] | 4.71 [3.82–8.21] | 5.67 [4.05–8.02] |
| RMS height, 50 µm grid (µm) | 7.27 [5.76–10.67] | 6.40 [5.30–11.64] | 7.43 [6.08–11.34] |
| Height P95–P5, 50 µm grid (µm) | 23.09 [17.99–35.48] | 18.39 [15.60–34.95] | 22.22 [17.11–35.48] |
| Median grid slope, 50 µm grid (°) | 0.97 [0.88–1.17] | 0.67 [0.62–0.85] | 0.67 [0.59–0.73] |
| Mean absolute height, 100 µm grid (µm) | 5.50 [4.18–8.35] | 4.62 [3.78–8.12] | 5.58 [3.99–7.94] |
| RMS height, 100 µm grid (µm) | 7.19 [5.70–10.58] | 6.24 [5.21–11.36] | 7.18 [5.90–11.20] |
| Height P95–P5, 100 µm grid (µm) | 22.90 [17.79–35.33] | 18.21 [15.38–34.33] | 22.09 [16.92–35.22] |
| Median grid slope, 100 µm grid (°) | 0.86 [0.77–1.02] | 0.59 [0.54–0.75] | 0.60 [0.52–0.64] |
| Mean absolute height, 200 µm grid (µm) | 5.33 [4.05–8.25] | 4.48 [3.69–7.95] | 5.37 [3.91–7.78] |
| RMS height, 200 µm grid (µm) | 6.99 [5.54–10.40] | 5.90 [4.81–10.88] | 7.00 [5.52–10.90] |
| Height P95–P5, 200 µm grid (µm) | 21.93 [17.17–34.44] | 17.90 [14.97–33.35] | 21.75 [16.65–34.67] |
| Median grid slope, 200 µm grid (°) | 0.65 [0.58–0.80] | 0.45 [0.41–0.60] | 0.46 [0.41–0.53] |

P05/P95 and SD refer to within-scan colour distributions; native edge and triangle percentiles are also calculated within scans. Table cells then summarise those scan-level values across specimens. Grid descriptors use common physical cell averaging; they do not equalise scanner resolution. All heights are plane-detrended; colours are uncalibrated.

# Table S2F Mature specimens TRIOS3

Values are medians [P25–P75] across 32 specimens per stage. Counts and triangle areas are rounded to whole units.

| Metric | Baseline | Partial | Clean |
| --- | --- | --- | --- |
| Mean L* | 48.23 [47.17–49.09] | 52.72 [51.85–53.37] | 54.34 [53.89–54.92] |
| Within-scan SD L* | 1.62 [1.37–1.88] | 1.38 [1.15–1.59] | 0.72 [0.66–0.81] |
| Median L* | 48.24 [47.20–49.17] | 53.16 [52.10–53.71] | 54.35 [53.88–54.92] |
| P05 L* | 45.48 [44.47–46.07] | 49.51 [48.15–51.06] | 53.15 [52.57–53.95] |
| P95 L* | 50.53 [49.89–51.65] | 54.12 [53.43–55.11] | 55.61 [55.20–55.99] |
| Mean a* | 46.33 [45.29–47.42] | 42.13 [40.22–43.47] | 36.87 [36.28–38.14] |
| Within-scan SD a* | 1.13 [1.00–1.23] | 3.41 [2.79–3.77] | 1.59 [1.22–1.71] |
| Median a* | 46.16 [45.25–47.50] | 42.14 [39.57–43.89] | 36.96 [36.07–38.50] |
| P05 a* | 44.65 [43.47–45.38] | 36.88 [35.56–37.96] | 34.41 [33.81–35.37] |
| P95 a* | 48.09 [46.75–49.49] | 48.50 [46.47–49.15] | 39.25 [38.44–40.37] |
| Mean b* | 21.03 [19.43–22.28] | 18.53 [18.22–18.97] | 17.79 [17.31–18.24] |
| Within-scan SD b* | 2.40 [2.21–2.73] | 1.86 [1.55–2.54] | 1.50 [1.38–1.63] |
| Median b* | 20.84 [19.32–22.19] | 18.51 [18.15–18.97] | 18.12 [17.22–18.55] |
| P05 b* | 17.82 [15.64–18.32] | 15.28 [14.77–15.66] | 15.09 [14.06–15.67] |
| P95 b* | 25.11 [23.42–26.58] | 21.81 [20.58–24.31] | 19.68 [19.41–19.84] |
| Encoded vertices | 610 [582–636] | 430 [411–469] | 428 [412–444] |
| Unique geometric vertices | 610 [582–636] | 430 [411–469] | 428 [412–444] |
| Triangle count | 1109 [1056–1155] | 764 [727–829] | 754 [721–776] |
| Mesh surface area (mm²) | 110.88 [110.50–111.28] | 110.63 [110.14–111.22] | 110.94 [110.32–111.34] |
| Encoded vertex density (mm⁻²) | 5.49 [5.27–5.70] | 3.90 [3.72–4.23] | 3.86 [3.75–4.00] |
| Geometric vertex density (mm⁻²) | 5.49 [5.27–5.70] | 3.90 [3.72–4.23] | 3.86 [3.75–4.00] |
| Triangle density (mm⁻²) | 9.99 [9.56–10.34] | 6.92 [6.59–7.46] | 6.78 [6.54–6.99] |
| Triangle area P10 (µm²) | 25317 [18305–31426] | 21644 [14591–25849] | 16853 [13983–23563] |
| Triangle area median (µm²) | 97683 [94703–104024] | 136372 [120918–145062] | 140601 [130839–147876] |
| Triangle area P90 (µm²) | 173037 [160330–183460] | 267108 [249840–291917] | 289959 [273006–310180] |
| Edge length P10 (µm) | 273.8 [258.4–299.9] | 267.4 [229.4–286.1] | 232.6 [208.7–275.0] |
| Edge length median (µm) | 498.8 [484.9–507.4] | 578.5 [556.7–596.4] | 583.7 [566.9–598.6] |
| Edge length P90 (µm) | 724.5 [702.4–744.8] | 906.4 [865.3–933.9] | 931.3 [914.2–963.4] |
| Native mean absolute height (µm) | 3.39 [2.88–4.38] | 3.02 [2.47–3.72] | 2.96 [2.33–4.60] |
| Native RMS height (µm) | 4.45 [3.70–6.36] | 4.09 [3.25–5.01] | 3.83 [3.16–6.00] |
| Native height P95–P5 (µm) | 13.93 [11.98–18.08] | 12.33 [9.67–14.95] | 11.95 [9.43–18.25] |
| Native face slope median (°) | 0.57 [0.52–0.62] | 0.28 [0.23–0.32] | 0.23 [0.21–0.27] |
| Native face slope P95 (°) | 1.95 [1.66–2.21] | 0.95 [0.82–1.04] | 0.68 [0.63–0.91] |
| Mean absolute height, 50 µm grid (µm) | 3.34 [2.86–4.29] | 3.01 [2.44–3.72] | 2.94 [2.31–4.60] |
| RMS height, 50 µm grid (µm) | 4.42 [3.64–6.29] | 4.07 [3.19–4.68] | 3.71 [3.09–5.81] |
| Height P95–P5, 50 µm grid (µm) | 13.82 [11.86–17.74] | 12.25 [9.61–14.86] | 11.90 [9.39–18.20] |
| Median grid slope, 50 µm grid (°) | 0.51 [0.47–0.56] | 0.26 [0.22–0.30] | 0.22 [0.20–0.26] |
| Mean absolute height, 100 µm grid (µm) | 3.30 [2.84–4.25] | 3.00 [2.42–3.71] | 2.92 [2.30–4.57] |
| RMS height, 100 µm grid (µm) | 4.39 [3.61–6.24] | 3.98 [3.13–4.64] | 3.61 [3.01–5.74] |
| Height P95–P5, 100 µm grid (µm) | 13.72 [11.76–17.63] | 12.21 [9.53–14.77] | 11.84 [9.32–18.04] |
| Median grid slope, 100 µm grid (°) | 0.46 [0.42–0.50] | 0.24 [0.21–0.28] | 0.21 [0.19–0.25] |
| Mean absolute height, 200 µm grid (µm) | 3.21 [2.78–4.06] | 2.97 [2.36–3.70] | 2.90 [2.28–4.53] |
| RMS height, 200 µm grid (µm) | 4.12 [3.52–5.67] | 3.73 [3.06–4.56] | 3.59 [2.98–5.69] |
| Height P95–P5, 200 µm grid (µm) | 13.31 [11.55–17.38] | 12.00 [9.33–14.64] | 11.73 [9.27–17.84] |
| Median grid slope, 200 µm grid (°) | 0.36 [0.34–0.41] | 0.22 [0.19–0.25] | 0.20 [0.17–0.23] |

P05/P95 and SD refer to within-scan colour distributions; native edge and triangle percentiles are also calculated within scans. Table cells then summarise those scan-level values across specimens. Grid descriptors use common physical cell averaging; they do not equalise scanner resolution. All heights are plane-detrended; colours are uncalibrated.

# Table S2G Mature specimens TRIOS5

Values are medians [P25–P75] across 32 specimens per stage. Counts and triangle areas are rounded to whole units.

| Metric | Baseline | Partial | Clean |
| --- | --- | --- | --- |
| Mean L* | 48.46 [46.69–49.47] | 53.76 [52.58–55.02] | 56.21 [55.67–56.70] |
| Within-scan SD L* | 2.14 [1.75–2.54] | 1.86 [1.56–2.52] | 0.76 [0.71–0.85] |
| Median L* | 48.55 [46.64–49.62] | 54.21 [52.91–55.18] | 56.28 [55.70–56.86] |
| P05 L* | 44.69 [43.30–45.77] | 49.30 [48.02–51.59] | 54.91 [54.37–55.23] |
| P95 L* | 51.88 [50.17–53.02] | 56.13 [55.05–56.77] | 57.39 [56.86–57.84] |
| Mean a* | 42.06 [40.00–42.60] | 33.90 [31.70–35.19] | 28.13 [27.39–29.12] |
| Within-scan SD a* | 1.73 [1.45–2.06] | 3.61 [3.03–4.39] | 1.48 [1.24–1.63] |
| Median a* | 42.16 [39.97–42.63] | 33.47 [30.59–34.95] | 28.40 [27.63–29.22] |
| P05 a* | 38.87 [37.21–40.14] | 28.44 [27.05–29.76] | 25.33 [24.58–26.32] |
| P95 a* | 44.35 [43.65–45.01] | 40.85 [38.80–43.01] | 30.25 [29.23–30.93] |
| Mean b* | 13.05 [12.23–14.04] | 11.30 [10.73–11.61] | 10.77 [10.32–11.43] |
| Within-scan SD b* | 1.58 [1.31–2.08] | 1.40 [1.14–1.71] | 0.98 [0.74–1.29] |
| Median b* | 12.83 [12.01–13.95] | 10.88 [10.64–11.55] | 10.59 [10.40–11.12] |
| P05 b* | 10.98 [9.74–11.81] | 9.22 [8.74–9.65] | 9.64 [9.32–9.72] |
| P95 b* | 16.01 [14.67–17.67] | 13.72 [12.40–15.18] | 12.05 [11.51–13.66] |
| Encoded vertices | 580 [563–642] | 364 [348–399] | 342 [328–368] |
| Unique geometric vertices | 580 [563–642] | 364 [348–399] | 342 [328–368] |
| Triangle count | 1056 [1024–1182] | 646 [613–711] | 601 [571–644] |
| Mesh surface area (mm²) | 104.45 [104.05–104.82] | 104.17 [103.62–104.92] | 104.04 [103.29–104.93] |
| Encoded vertex density (mm⁻²) | 5.56 [5.39–6.15] | 3.50 [3.33–3.84] | 3.30 [3.15–3.52] |
| Geometric vertex density (mm⁻²) | 5.56 [5.39–6.15] | 3.50 [3.33–3.84] | 3.30 [3.15–3.52] |
| Triangle density (mm⁻²) | 10.10 [9.78–11.34] | 6.16 [5.87–6.83] | 5.81 [5.49–6.22] |
| Triangle area P10 (µm²) | 45119 [41590–49456] | 68066 [54294–76870] | 72705 [41055–84995] |
| Triangle area median (µm²) | 93601 [83406–97862] | 153304 [140301–163138] | 163383 [155834–174155] |
| Triangle area P90 (µm²) | 157482 [141126–164163] | 268639 [239991–280659] | 278429 [268335–288498] |
| Edge length P10 (µm) | 315.7 [298.5–327.1] | 387.4 [346.4–405.0] | 399.8 [343.0–417.4] |
| Edge length median (µm) | 485.7 [464.2–499.2] | 615.9 [588.0–634.9] | 645.8 [619.8–655.0] |
| Edge length P90 (µm) | 687.7 [660.1–702.1] | 902.1 [860.7–948.5] | 929.7 [910.0–946.3] |
| Native mean absolute height (µm) | 3.66 [2.88–4.24] | 2.84 [2.30–3.57] | 3.29 [2.37–4.01] |
| Native RMS height (µm) | 4.80 [3.68–5.55] | 3.55 [2.93–4.49] | 4.21 [3.01–5.03] |
| Native height P95–P5 (µm) | 15.67 [12.10–18.18] | 11.79 [9.50–14.94] | 14.14 [9.92–15.95] |
| Native face slope median (°) | 0.75 [0.65–0.88] | 0.29 [0.26–0.32] | 0.28 [0.24–0.33] |
| Native face slope P95 (°) | 2.70 [2.26–3.24] | 0.85 [0.76–1.01] | 0.75 [0.65–0.82] |
| Mean absolute height, 50 µm grid (µm) | 3.65 [2.87–4.22] | 2.84 [2.29–3.57] | 3.26 [2.37–3.98] |
| RMS height, 50 µm grid (µm) | 4.74 [3.66–5.52] | 3.54 [2.91–4.48] | 4.19 [3.00–5.00] |
| Height P95–P5, 50 µm grid (µm) | 15.63 [12.09–18.02] | 11.75 [9.50–14.90] | 14.05 [9.87–15.50] |
| Median grid slope, 50 µm grid (°) | 0.67 [0.59–0.78] | 0.28 [0.25–0.31] | 0.27 [0.23–0.31] |
| Mean absolute height, 100 µm grid (µm) | 3.61 [2.82–4.19] | 2.83 [2.27–3.57] | 3.25 [2.36–3.95] |
| RMS height, 100 µm grid (µm) | 4.67 [3.62–5.48] | 3.53 [2.88–4.46] | 4.17 [2.98–4.96] |
| Height P95–P5, 100 µm grid (µm) | 15.44 [11.91–17.85] | 11.68 [9.42–14.82] | 14.00 [9.82–15.40] |
| Median grid slope, 100 µm grid (°) | 0.60 [0.52–0.69] | 0.26 [0.24–0.29] | 0.25 [0.21–0.29] |
| Mean absolute height, 200 µm grid (µm) | 3.50 [2.68–4.10] | 2.79 [2.23–3.55] | 3.21 [2.32–3.90] |
| RMS height, 200 µm grid (µm) | 4.54 [3.45–5.37] | 3.49 [2.82–4.43] | 4.12 [2.94–4.92] |
| Height P95–P5, 200 µm grid (µm) | 15.03 [11.32–17.54] | 11.54 [9.25–14.58] | 13.85 [9.56–15.25] |
| Median grid slope, 200 µm grid (°) | 0.46 [0.41–0.53] | 0.23 [0.21–0.26] | 0.22 [0.19–0.25] |

P05/P95 and SD refer to within-scan colour distributions; native edge and triangle percentiles are also calculated within scans. Table cells then summarise those scan-level values across specimens. Grid descriptors use common physical cell averaging; they do not equalise scanner resolution. All heights are plane-detrended; colours are uncalibrated.

# Table S2H Mature specimens LABscanner

Values are medians [P25–P75] across 32 specimens per stage. Counts and triangle areas are rounded to whole units.

| Metric | Baseline | Partial | Clean |
| --- | --- | --- | --- |
| Mean L* | 55.73 [55.21–57.66] | 54.37 [53.84–54.79] | 56.11 [55.15–58.01] |
| Within-scan SD L* | 2.11 [1.66–2.36] | 1.30 [1.19–1.39] | 1.51 [1.36–1.79] |
| Median L* | 55.45 [55.06–57.49] | 54.36 [53.87–54.98] | 56.11 [55.15–58.17] |
| P05 L* | 53.07 [52.29–54.40] | 51.97 [51.81–52.54] | 53.40 [52.69–55.57] |
| P95 L* | 59.66 [58.23–61.55] | 56.47 [55.86–57.09] | 58.41 [57.77–61.39] |
| Mean a* | 34.77 [32.51–35.51] | 32.65 [31.02–34.04] | 29.94 [26.22–30.51] |
| Within-scan SD a* | 3.06 [2.65–3.30] | 3.08 [2.81–3.60] | 2.18 [1.98–2.55] |
| Median a* | 35.41 [32.70–35.99] | 32.67 [31.15–34.04] | 30.09 [26.82–31.12] |
| P05 a* | 28.96 [27.19–30.35] | 27.18 [25.70–28.17] | 25.79 [22.12–26.83] |
| P95 a* | 39.04 [36.80–39.42] | 37.86 [36.48–39.29] | 32.40 [29.21–33.02] |
| Mean b* | 26.84 [25.13–27.48] | 23.94 [23.17–24.88] | 23.83 [22.97–24.99] |
| Within-scan SD b* | 2.74 [2.44–3.00] | 2.58 [2.32–2.78] | 2.38 [1.93–2.64] |
| Median b* | 26.77 [25.71–27.21] | 24.53 [23.12–25.77] | 24.38 [23.03–26.00] |
| P05 b* | 21.95 [20.91–22.90] | 19.24 [18.94–20.21] | 19.58 [18.80–20.11] |
| P95 b* | 31.14 [29.42–32.09] | 27.78 [26.85–28.58] | 27.15 [26.21–28.32] |
| Encoded vertices | 8340 [7723–8800] | 4909 [4702–5226] | 4576 [4484–4729] |
| Unique geometric vertices | 8340 [7723–8800] | 4909 [4702–5226] | 4576 [4484–4729] |
| Triangle count | 16326 [15098–17244] | 9550 [9129–10169] | 8888 [8694–9199] |
| Mesh surface area (mm²) | 106.04 [105.95–106.15] | 105.07 [104.96–105.15] | 104.88 [104.79–105.07] |
| Encoded vertex density (mm⁻²) | 78.62 [72.86–82.96] | 46.66 [44.72–49.73] | 43.60 [42.80–45.02] |
| Geometric vertex density (mm⁻²) | 78.62 [72.86–82.96] | 46.66 [44.72–49.73] | 43.60 [42.80–45.02] |
| Triangle density (mm⁻²) | 153.91 [142.42–162.55] | 90.78 [86.81–96.77] | 84.69 [83.06–87.57] |
| Triangle area P10 (µm²) | 3226 [3210–3268] | 5131 [4578–5583] | 6176 [5844–6292] |
| Triangle area median (µm²) | 5801 [5512–6228] | 10218 [9463–10696] | 11010 [10656–11289] |
| Triangle area P90 (µm²) | 10878 [10314–11892] | 17804 [16638–18345] | 18383 [17960–18731] |
| Edge length P10 (µm) | 80.3 [80.1–80.8] | 102.4 [96.2–107.2] | 112.9 [109.0–113.4] |
| Edge length median (µm) | 123.9 [118.6–131.4] | 165.2 [160.3–168.5] | 172.0 [168.8–173.4] |
| Edge length P90 (µm) | 192.0 [187.1–199.8] | 245.3 [239.3–252.6] | 252.8 [250.0–256.7] |
| Native mean absolute height (µm) | 5.07 [4.66–5.57] | 4.31 [3.93–5.17] | 3.73 [3.01–4.37] |
| Native RMS height (µm) | 6.60 [6.03–7.08] | 5.56 [5.26–6.80] | 4.73 [3.84–5.77] |
| Native height P95–P5 (µm) | 21.16 [19.47–23.22] | 17.96 [17.04–21.12] | 15.41 [12.47–18.41] |
| Native face slope median (°) | 3.73 [3.28–4.05] | 1.86 [1.77–2.03] | 1.79 [1.71–1.82] |
| Native face slope P95 (°) | 10.25 [8.87–11.17] | 4.57 [4.30–5.37] | 3.70 [3.59–3.88] |
| Mean absolute height, 50 µm grid (µm) | 4.96 [4.55–5.41] | 4.29 [3.90–5.09] | 3.68 [2.99–4.34] |
| RMS height, 50 µm grid (µm) | 6.44 [5.87–6.93] | 5.51 [5.23–6.74] | 4.67 [3.80–5.72] |
| Height P95–P5, 50 µm grid (µm) | 20.63 [19.08–22.72] | 17.84 [16.89–20.89] | 15.20 [12.38–18.12] |
| Median grid slope, 50 µm grid (°) | 2.88 [2.53–3.15] | 1.42 [1.37–1.56] | 1.36 [1.30–1.39] |
| Mean absolute height, 100 µm grid (µm) | 4.76 [4.27–5.18] | 4.23 [3.85–5.01] | 3.62 [2.93–4.29] |
| RMS height, 100 µm grid (µm) | 6.07 [5.50–6.60] | 5.43 [5.12–6.62] | 4.58 [3.72–5.65] |
| Height P95–P5, 100 µm grid (µm) | 19.55 [17.94–21.52] | 17.61 [16.60–20.59] | 14.92 [12.09–17.91] |
| Median grid slope, 100 µm grid (°) | 1.97 [1.75–2.16] | 1.05 [1.01–1.15] | 0.97 [0.93–1.00] |
| Mean absolute height, 200 µm grid (µm) | 3.96 [3.73–4.51] | 4.09 [3.69–4.83] | 3.47 [2.80–4.16] |
| RMS height, 200 µm grid (µm) | 5.09 [4.77–5.76] | 5.20 [4.85–6.32] | 4.34 [3.54–5.45] |
| Height P95–P5, 200 µm grid (µm) | 16.56 [15.68–18.80] | 17.05 [15.72–19.81] | 14.18 [11.59–17.41] |
| Median grid slope, 200 µm grid (°) | 0.82 [0.74–0.87] | 0.59 [0.55–0.62] | 0.51 [0.50–0.54] |

P05/P95 and SD refer to within-scan colour distributions; native edge and triangle percentiles are also calculated within scans. Table cells then summarise those scan-level values across specimens. Grid descriptors use common physical cell averaging; they do not equalise scanner resolution. All heights are plane-detrended; colours are uncalibrated.