# Paired colour and geometric comparisons

Table B9. Within-specimen colour and geometric changes across experimental stages. Panel A.

Fresh specimens · Baseline to clean · n = 16 matched specimens per scanner

| Metric | iTERO | TRIOS3 | TRIOS5 | LABscanner |
| --- | --- | --- | --- | --- |
| Colour difference ΔE₀₀ | 8.70 [7.10, 9.95] | 6.12 [5.76, 6.34] | 7.66 [6.20, 8.72] | 2.94 [2.48, 3.47] |
| Δ mean L* | 7.61 [5.96, 9.33] | 4.91 [4.81, 5.06] | 6.21 [3.97, 6.98] | 0.53 [−0.38, 1.53] |
| Δ mean a* | −10.74 [−11.70, −8.05] | −9.43 [−10.37, −9.11] | −10.86 [−11.96, −9.81] | −5.45 [−6.24, −5.04] |
| Δ mean b* | −4.72 [−6.19, −3.81] | −4.79 [−6.11, −3.21] | −5.00 [−6.41, −3.57] | −1.80 [−2.75, 0.68] |
| Δ RMS height 50 µm grid (µm) | 1.49 [−0.39, 7.06] | −0.24 [−1.14, 0.51] | −0.52 [−1.63, 0.50] | −1.67 [−2.63, −1.18] |
| Δ RMS height 100 µm grid (µm) | 1.51 [−0.39, 7.07] | −0.23 [−1.14, 0.49] | −0.50 [−1.58, 0.58] | −1.35 [−2.43, −0.85] |
| Δ RMS height 200 µm grid (µm) | 1.43 [−0.40, 6.83] | −0.18 [−0.96, 0.50] | −0.47 [−1.45, 0.76] | −0.49 [−1.75, −0.15] |
| Δ native mean absolute height (µm) | 0.97 [−0.29, 5.70] | −0.19 [−0.73, 0.38] | −0.35 [−1.24, 0.65] | −1.23 [−2.04, −0.96] |
| Δ native RMS height (µm) | 1.52 [−0.34, 7.08] | −0.25 [−1.16, 0.54] | −0.53 [−1.66, 0.45] | −1.81 [−2.72, −1.32] |
| Δ native height P95−P5 (µm) | 4.20 [−1.92, 20.35] | −0.82 [−3.55, 2.43] | −0.74 [−4.81, 4.07] | −5.66 [−8.69, −3.76] |
| Δ median native face slope (°) | −0.15 [−0.31, −0.07] | −0.29 [−0.36, −0.20] | −0.31 [−0.44, −0.17] | −2.29 [−2.47, −1.98] |
| Δ native face slope P95 (°) | −0.38 [−1.18, 0.67] | −1.00 [−1.33, −0.74] | −1.16 [−2.35, −0.87] | −7.02 [−7.54, −6.64] |
| Δ triangle density (mm⁻²) | −9.3 [−12.9, −5.0] | −3.3 [−4.2, −2.5] | −3.5 [−4.4, −2.4] | −71.1 [−78.2, −60.2] |
| Triangle density ratio later / earlier | 0.73 [0.65, 0.86] | 0.65 [0.59, 0.73] | 0.62 [0.56, 0.71] | 0.55 [0.53, 0.59] |
| Δ median edge length (µm) | 46.3 [23.6, 63.8] | 96.1 [73.6, 143.1] | 144.3 [95.5, 163.4] | 47.0 [41.6, 51.7] |
| Δ median triangle area (µm²) | 10044 [5170, 13196] | 50222 [38763, 73975] | 63851 [49384, 76341] | 5130 [4654, 5475] |

Values are medians [95% percentile bootstrap confidence intervals] from 1,000 specimen-level resamples. Signed changes are later minus earlier; the triangle-density ratio is later divided by earlier. A ratio below 1 indicates a reduction. ΔE₀₀ compares area-weighted scan-mean CIELAB colours and is unsigned.

Heights and slopes are relative to a fitted plane. Grid widths of 50, 100 and 200 µm specify physical cell averaging. Native descriptors use the exported mesh. Colour is uncalibrated; topography is scanner-derived. Negative values that round to zero are shown as zero.

Table B9. Within-specimen colour and geometric changes across experimental stages. Panel B.

Mature specimens · Baseline to partial cleaning · n = 32 matched specimens per scanner

| Metric | iTERO | TRIOS3 | TRIOS5 | LABscanner |
| --- | --- | --- | --- | --- |
| Colour difference ΔE₀₀ | 5.67 [5.00, 6.49] | 4.95 [4.43, 5.83] | 6.35 [5.63, 6.87] | 2.71 [2.14, 4.27] |
| Δ mean L* | 4.23 [3.82, 5.26] | 4.37 [3.68, 5.48] | 5.47 [4.59, 6.02] | −1.58 [−2.85, −0.86] |
| Δ mean a* | −9.48 [−11.12, −6.22] | −3.95 [−5.29, −3.53] | −7.90 [−8.93, −7.05] | −2.23 [−3.71, 0.01] |
| Δ mean b* | −1.64 [−2.61, −0.94] | −2.62 [−3.44, −1.53] | −1.96 [−2.60, −1.01] | −2.70 [−3.95, −0.88] |
| Δ RMS height 50 µm grid (µm) | −1.13 [−2.45, −0.38] | −0.53 [−0.92, −0.03] | −1.05 [−1.47, −0.68] | −0.49 [−0.79, 0.00] |
| Δ RMS height 100 µm grid (µm) | −1.14 [−2.28, −0.59] | −0.47 [−0.90, −0.02] | −1.02 [−1.43, −0.66] | −0.30 [−0.59, 0.16] |
| Δ RMS height 200 µm grid (µm) | −1.07 [−2.21, −0.52] | −0.48 [−0.82, 0.02] | −0.95 [−1.34, −0.58] | 0.34 [−0.18, 0.61] |
| Δ native mean absolute height (µm) | −0.69 [−1.30, −0.05] | −0.36 [−0.59, −0.08] | −0.79 [−1.08, −0.43] | −0.51 [−0.83, −0.06] |
| Δ native RMS height (µm) | −1.12 [−2.38, −0.21] | −0.57 [−0.98, −0.03] | −1.07 [−1.49, −0.69] | −0.54 [−0.90, −0.07] |
| Δ native height P95−P5 (µm) | −3.13 [−5.30, −0.63] | −1.93 [−2.85, −0.08] | −3.36 [−4.43, −2.40] | −2.03 [−3.18, −0.20] |
| Δ median native face slope (°) | −0.35 [−0.46, −0.23] | −0.26 [−0.33, −0.24] | −0.45 [−0.55, −0.39] | −1.80 [−2.09, −1.50] |
| Δ native face slope P95 (°) | −1.09 [−1.44, −0.55] | −0.83 [−1.18, −0.69] | −1.84 [−2.20, −1.47] | −5.21 [−5.87, −4.16] |
| Δ triangle density (mm⁻²) | −14.8 [−18.0, −11.1] | −2.8 [−3.3, −2.6] | −3.8 [−4.5, −3.3] | −61.0 [−69.9, −54.8] |
| Triangle density ratio later / earlier | 0.64 [0.60, 0.72] | 0.72 [0.67, 0.73] | 0.63 [0.57, 0.66] | 0.60 [0.56, 0.64] |
| Δ median edge length (µm) | 49.0 [41.4, 73.1] | 85.6 [69.1, 96.3] | 129.5 [112.1, 139.3] | 39.0 [35.4, 45.7] |
| Δ median triangle area (µm²) | 10005 [8040, 15358] | 38662 [30520, 43599] | 56762 [51036, 68124] | 4415 [3712, 4680] |

Values are medians [95% percentile bootstrap confidence intervals] from 1,000 specimen-level resamples. Signed changes are later minus earlier; the triangle-density ratio is later divided by earlier. A ratio below 1 indicates a reduction. ΔE₀₀ compares area-weighted scan-mean CIELAB colours and is unsigned.

Heights and slopes are relative to a fitted plane. Grid widths of 50, 100 and 200 µm specify physical cell averaging. Native descriptors use the exported mesh. Colour is uncalibrated; topography is scanner-derived. Negative values that round to zero are shown as zero.

Table B9. Within-specimen colour and geometric changes across experimental stages. Panel C.

Mature specimens · Partial cleaning to clean · n = 32 matched specimens per scanner

| Metric | iTERO | TRIOS3 | TRIOS5 | LABscanner |
| --- | --- | --- | --- | --- |
| Colour difference ΔE₀₀ | 3.39 [2.77, 4.53] | 2.66 [1.86, 3.09] | 3.42 [2.41, 3.78] | 2.16 [1.70, 3.29] |
| Δ mean L* | 2.85 [1.86, 3.89] | 1.90 [1.17, 2.49] | 2.62 [1.55, 3.22] | 1.38 [0.92, 1.94] |
| Δ mean a* | −4.52 [−5.51, −3.68] | −4.22 [−5.04, −3.58] | −4.71 [−6.14, −3.74] | −2.72 [−5.11, −1.88] |
| Δ mean b* | −0.26 [−0.73, 0.52] | −0.63 [−1.21, −0.46] | −0.43 [−0.67, −0.15] | 0.02 [−0.72, 0.63] |
| Δ RMS height 50 µm grid (µm) | −0.04 [−0.93, 1.87] | −0.02 [−0.28, 0.53] | 0.02 [−0.30, 0.58] | −1.24 [−1.54, −0.61] |
| Δ RMS height 100 µm grid (µm) | 0.01 [−0.84, 1.67] | −0.04 [−0.31, 0.39] | 0.02 [−0.29, 0.58] | −1.23 [−1.49, −0.61] |
| Δ RMS height 200 µm grid (µm) | −0.14 [−0.81, 1.54] | −0.05 [−0.30, 0.40] | 0.03 [−0.29, 0.57] | −1.18 [−1.48, −0.53] |
| Δ native mean absolute height (µm) | −0.07 [−0.60, 1.08] | −0.02 [−0.25, 0.30] | 0.04 [−0.27, 0.42] | −0.88 [−1.19, −0.39] |
| Δ native RMS height (µm) | −0.02 [−0.92, 1.89] | 0.01 [−0.25, 0.54] | 0.01 [−0.29, 0.59] | −1.25 [−1.55, −0.59] |
| Δ native height P95−P5 (µm) | −0.92 [−2.54, 4.84] | −0.40 [−0.86, 1.18] | −0.18 [−0.83, 1.74] | −3.99 [−4.65, −2.33] |
| Δ median native face slope (°) | −0.05 [−0.14, 0.02] | −0.04 [−0.06, −0.01] | −0.02 [−0.04, 0.01] | −0.16 [−0.26, −0.03] |
| Δ native face slope P95 (°) | −0.06 [−0.35, 0.16] | −0.24 [−0.32, −0.08] | −0.08 [−0.15, −0.01] | −1.02 [−1.32, −0.73] |
| Δ triangle density (mm⁻²) | 1.0 [−3.2, 2.6] | −0.3 [−0.6, 0.2] | −0.4 [−0.6, 0.1] | −7.4 [−12.3, −4.3] |
| Triangle density ratio later / earlier | 1.03 [0.88, 1.09] | 0.96 [0.91, 1.03] | 0.93 [0.89, 1.03] | 0.92 [0.88, 0.95] |
| Δ median edge length (µm) | −6.8 [−14.2, 17.9] | 11.6 [−12.8, 21.3] | 18.1 [−3.6, 38.7] | 7.3 [4.7, 10.7] |
| Δ median triangle area (µm²) | −911 [−4014, 4813] | 6357 [695, 14345] | 8872 [−296, 19531] | 988 [592, 1477] |

Values are medians [95% percentile bootstrap confidence intervals] from 1,000 specimen-level resamples. Signed changes are later minus earlier; the triangle-density ratio is later divided by earlier. A ratio below 1 indicates a reduction. ΔE₀₀ compares area-weighted scan-mean CIELAB colours and is unsigned.

Heights and slopes are relative to a fitted plane. Grid widths of 50, 100 and 200 µm specify physical cell averaging. Native descriptors use the exported mesh. Colour is uncalibrated; topography is scanner-derived. Negative values that round to zero are shown as zero.

Table B9. Within-specimen colour and geometric changes across experimental stages. Panel D.

Mature specimens · Baseline to clean · n = 32 matched specimens per scanner

| Metric | iTERO | TRIOS3 | TRIOS5 | LABscanner |
| --- | --- | --- | --- | --- |
| Colour difference ΔE₀₀ | 8.24 [7.26, 9.80] | 6.88 [6.27, 7.37] | 9.34 [8.60, 9.88] | 2.93 [2.51, 3.18] |
| Δ mean L* | 7.23 [5.44, 8.90] | 5.96 [5.46, 6.72] | 7.88 [6.87, 8.63] | 0.06 [−0.46, 0.85] |
| Δ mean a* | −14.30 [−15.70, −12.21] | −8.73 [−10.12, −8.13] | −12.91 [−14.64, −12.16] | −5.43 [−6.14, −4.51] |
| Δ mean b* | −1.74 [−2.91, −0.72] | −3.52 [−4.01, −2.37] | −2.59 [−3.01, −1.47] | −2.76 [−3.77, −1.80] |
| Δ RMS height 50 µm grid (µm) | 0.12 [−1.94, 1.06] | −0.66 [−0.97, −0.22] | −0.84 [−1.27, −0.23] | −1.69 [−2.30, −1.17] |
| Δ RMS height 100 µm grid (µm) | −0.28 [−1.87, 1.06] | −0.66 [−1.03, −0.20] | −0.81 [−1.23, −0.16] | −1.43 [−2.01, −0.92] |
| Δ RMS height 200 µm grid (µm) | −0.34 [−1.76, 1.04] | −0.61 [−0.98, −0.23] | −0.67 [−1.06, −0.09] | −0.72 [−1.37, −0.37] |
| Δ native mean absolute height (µm) | −0.50 [−1.67, 0.76] | −0.53 [−0.77, −0.25] | −0.59 [−0.93, −0.14] | −1.25 [−1.85, −1.00] |
| Δ native RMS height (µm) | 0.10 [−2.23, 1.28] | −0.68 [−0.91, −0.50] | −0.86 [−1.27, −0.29] | −1.74 [−2.41, −1.32] |
| Δ native height P95−P5 (µm) | −1.99 [−6.91, 2.41] | −2.38 [−3.68, −1.22] | −2.78 [−3.90, −1.23] | −5.41 [−7.55, −4.00] |
| Δ median native face slope (°) | −0.36 [−0.50, −0.29] | −0.29 [−0.35, −0.25] | −0.49 [−0.58, −0.38] | −1.98 [−2.08, −1.72] |
| Δ native face slope P95 (°) | −0.91 [−1.61, −0.52] | −1.21 [−1.35, −0.94] | −1.87 [−2.23, −1.58] | −6.39 [−7.30, −5.74] |
| Δ triangle density (mm⁻²) | −14.3 [−17.9, −11.5] | −3.0 [−3.5, −2.8] | −4.4 [−4.8, −4.0] | −69.3 [−76.3, −61.5] |
| Triangle density ratio later / earlier | 0.66 [0.58, 0.70] | 0.69 [0.66, 0.71] | 0.57 [0.54, 0.62] | 0.55 [0.52, 0.59] |
| Δ median edge length (µm) | 54.0 [43.9, 71.8] | 87.7 [75.2, 100.0] | 156.3 [127.9, 171.2] | 48.1 [42.7, 52.7] |
| Δ median triangle area (µm²) | 11151 [8867, 14672] | 39886 [35667, 45305] | 69853 [65115, 76767] | 5126 [4859, 5523] |

Values are medians [95% percentile bootstrap confidence intervals] from 1,000 specimen-level resamples. Signed changes are later minus earlier; the triangle-density ratio is later divided by earlier. A ratio below 1 indicates a reduction. ΔE₀₀ compares area-weighted scan-mean CIELAB colours and is unsigned.

Heights and slopes are relative to a fitted plane. Grid widths of 50, 100 and 200 µm specify physical cell averaging. Native descriptors use the exported mesh. Colour is uncalibrated; topography is scanner-derived. Negative values that round to zero are shown as zero.

Table B10. Matched between-scanner colour and geometric differences. Panel A.

Fresh specimens · Baseline · n = 16 matched specimens per scanner pair

| Metric | iTERO → TRIOS3 | iTERO → TRIOS5 | iTERO → LABscanner | TRIOS3 → TRIOS5 | TRIOS3 → LABscanner | TRIOS5 → LABscanner |
| --- | --- | --- | --- | --- | --- | --- |
| Colour difference ΔE₀₀ | 4.61 [3.89, 6.55] | 4.87 [4.65, 7.81] | 12.69 [11.89, 14.06] | 3.73 [3.31, 4.99] | 8.48 [7.17, 9.32] | 8.64 [4.89, 9.16] |
| Δ mean L* | 1.47 [0.89, 4.27] | 3.27 [2.57, 5.58] | 7.25 [5.80, 10.05] | 1.37 [0.31, 2.80] | 5.56 [4.43, 6.59] | 3.71 [0.61, 5.89] |
| Δ mean a* | −4.51 [−5.79, −3.43] | −11.61 [−13.32, −9.82] | −15.11 [−17.19, −13.76] | −6.90 [−7.89, −5.92] | −10.34 [−11.57, −9.23] | −4.16 [−5.40, −1.08] |
| Δ mean b* | 5.39 [4.49, 7.15] | −0.40 [−0.99, 0.49] | 9.52 [7.65, 11.41] | −6.21 [−6.84, −5.23] | 4.10 [2.32, 5.36] | 10.64 [8.44, 11.13] |
| Wasserstein L* | 2.02 [1.46, 4.27] | 3.29 [2.82, 5.58] | 7.24 [5.82, 10.05] | 1.79 [1.26, 2.80] | 5.56 [4.42, 6.59] | 3.72 [1.61, 5.88] |
| Wasserstein a* | 4.49 [3.47, 5.77] | 11.59 [9.81, 13.33] | 15.15 [13.74, 17.18] | 6.92 [5.93, 7.86] | 10.32 [9.23, 11.65] | 4.20 [1.67, 5.40] |
| Wasserstein b* | 5.40 [4.47, 7.15] | 0.88 [0.73, 1.40] | 9.55 [7.65, 11.42] | 6.19 [5.24, 6.82] | 4.15 [2.34, 5.39] | 10.68 [8.47, 11.15] |
| Δ RMS height 50 µm grid (µm) | −1.45 [−1.94, −0.98] | −1.21 [−2.03, −0.79] | 0.66 [0.38, 1.38] | 0.06 [−0.12, 0.55] | 2.45 [1.64, 3.19] | 2.19 [1.57, 2.92] |
| Δ RMS height 100 µm grid (µm) | −1.38 [−1.86, −0.91] | −1.11 [−1.96, −0.57] | 0.53 [0.17, 1.17] | 0.06 [−0.13, 0.54] | 2.15 [1.34, 2.84] | 1.89 [1.36, 2.62] |
| Δ RMS height 200 µm grid (µm) | −1.35 [−1.73, −0.83] | −1.02 [−1.88, −0.46] | −0.02 [−0.49, 0.36] | 0.07 [−0.17, 0.49] | 1.33 [0.75, 1.90] | 0.91 [0.53, 1.67] |
| Δ native mean absolute height (µm) | −1.07 [−1.35, −0.37] | −0.92 [−1.56, −0.39] | 0.98 [0.41, 1.25] | 0.02 [−0.15, 0.43] | 2.01 [1.46, 2.56] | 1.90 [1.28, 2.42] |
| Δ native RMS height (µm) | −1.53 [−2.08, −1.07] | −1.36 [−2.11, −0.93] | 0.52 [0.29, 1.50] | 0.06 [−0.11, 0.56] | 2.56 [1.73, 3.32] | 2.33 [1.66, 3.03] |
| Δ native height P95−P5 (µm) | −4.17 [−6.40, −2.75] | −3.93 [−6.04, −1.28] | 4.19 [1.81, 5.71] | −0.05 [−0.93, 1.08] | 8.41 [6.16, 10.71] | 7.54 [5.46, 9.87] |
| Δ median native face slope (°) | −0.44 [−0.55, −0.37] | −0.38 [−0.43, −0.28] | 3.02 [2.70, 3.25] | 0.02 [−0.01, 0.20] | 3.46 [3.17, 3.65] | 3.20 [3.07, 3.65] |
| Δ native face slope P95 (°) | −1.16 [−1.44, −0.83] | −0.61 [−0.94, −0.10] | 8.01 [7.26, 8.44] | 0.28 [−0.06, 1.07] | 9.06 [8.35, 9.59] | 8.33 [7.73, 8.89] |
| Δ triangle density (mm⁻²) | −25.9 [−28.1, −24.1] | −26.3 [−28.2, −24.4] | 119.0 [108.9, 130.7] | −0.3 [−0.8, 0.3] | 146.3 [132.8, 158.1] | 145.2 [133.4, 159.0] |
| Δ median edge length (µm) | 238.5 [236.5, 263.0] | 247.0 [227.4, 260.9] | −128.6 [−132.9, −118.8] | 1.7 [−16.3, 13.0] | −370.4 [−394.2, −358.2] | −371.5 [−393.8, −346.6] |
| Δ median triangle area (µm²) | 75273 [68999, 82005] | 73171 [66507, 81871] | −19163 [−19917, −17953] | 1574 [−8672, 6226] | −93886 [−102723, −87047] | −91658 [−101569, −85587] |

Values are medians [95% percentile bootstrap confidence intervals] from 1,000 specimen-level resamples. Column labels A → B indicate signed differences B minus A for the same specimen and stage. ΔE₀₀ and Wasserstein distances are unsigned; Wasserstein distances use the indicated channel units.

ΔE₀₀ compares scan-mean CIELAB vectors. Wasserstein distances compare separate area-weighted L*, a* and b* distributions. Geometric comparisons use scan-level summaries without assuming pointwise correspondence. Colours are uncalibrated, and denser sampling does not establish greater geometric accuracy.

Table B10. Matched between-scanner colour and geometric differences. Panel B.

Fresh specimens · Clean · n = 16 matched specimens per scanner pair

| Metric | iTERO → TRIOS3 | iTERO → TRIOS5 | iTERO → LABscanner | TRIOS3 → TRIOS5 | TRIOS3 → LABscanner | TRIOS5 → LABscanner |
| --- | --- | --- | --- | --- | --- | --- |
| Colour difference ΔE₀₀ | 4.85 [4.33, 5.08] | 6.00 [5.83, 6.15] | 11.33 [10.53, 11.59] | 5.09 [5.01, 5.15] | 6.61 [6.13, 6.93] | 8.01 [7.73, 8.28] |
| Δ mean L* | −0.44 [−1.48, 0.18] | 2.35 [1.37, 3.01] | 0.46 [−0.15, 1.31] | 2.83 [2.47, 3.10] | 1.33 [0.59, 1.87] | −1.72 [−2.34, −1.10] |
| Δ mean a* | −4.39 [−4.81, −3.49] | −13.01 [−13.59, −12.24] | −10.49 [−11.79, −9.26] | −8.80 [−8.84, −8.28] | −6.31 [−6.58, −5.39] | 2.54 [1.25, 3.06] |
| Δ mean b* | 6.02 [5.53, 6.29] | −0.36 [−0.88, −0.05] | 12.61 [11.72, 14.06] | −6.33 [−6.68, −6.15] | 6.66 [5.99, 7.85] | 13.03 [12.41, 13.96] |
| Wasserstein L* | 1.80 [1.64, 2.03] | 2.38 [1.75, 3.01] | 1.11 [0.76, 1.38] | 2.83 [2.49, 3.09] | 1.42 [1.11, 1.98] | 1.90 [1.20, 2.38] |
| Wasserstein a* | 4.43 [3.61, 4.87] | 13.04 [12.26, 13.59] | 10.40 [9.30, 11.77] | 8.78 [8.26, 8.87] | 6.28 [5.34, 6.61] | 2.60 [1.42, 3.09] |
| Wasserstein b* | 6.06 [5.48, 6.35] | 0.77 [0.51, 1.14] | 12.68 [11.78, 14.15] | 6.47 [6.26, 6.71] | 6.66 [5.93, 7.85] | 13.10 [12.45, 14.08] |
| Δ RMS height 50 µm grid (µm) | −2.18 [−7.57, −1.46] | −2.29 [−7.54, −1.47] | −2.92 [−9.16, −0.20] | −0.41 [−0.80, −0.05] | 1.11 [0.22, 1.66] | 1.45 [0.02, 2.16] |
| Δ RMS height 100 µm grid (µm) | −2.05 [−7.40, −1.35] | −2.04 [−7.40, −1.33] | −2.81 [−9.01, −0.13] | −0.41 [−0.78, −0.04] | 1.07 [0.17, 1.57] | 1.38 [−0.06, 2.11] |
| Δ RMS height 200 µm grid (µm) | −1.93 [−7.17, −1.21] | −1.83 [−7.15, −1.25] | −2.67 [−8.82, −0.15] | −0.41 [−0.76, −0.04] | 0.94 [0.01, 1.33] | 1.21 [−0.27, 1.99] |
| Δ native mean absolute height (µm) | −1.36 [−5.57, −0.94] | −1.46 [−5.65, −0.81] | −1.74 [−6.73, 0.07] | −0.30 [−0.73, 0.14] | 0.89 [0.11, 1.23] | 0.99 [−0.12, 1.69] |
| Δ native RMS height (µm) | −2.36 [−7.69, −1.62] | −2.65 [−7.66, −1.61] | −3.17 [−9.32, −0.35] | −0.41 [−0.84, −0.05] | 1.12 [0.21, 1.71] | 1.49 [0.06, 2.18] |
| Δ native height P95−P5 (µm) | −7.12 [−20.79, −3.94] | −5.83 [−21.39, −3.51] | −6.66 [−24.64, 0.58] | −1.15 [−2.16, −0.59] | 3.95 [0.15, 5.72] | 5.27 [−0.85, 7.52] |
| Δ median native face slope (°) | −0.53 [−0.57, −0.47] | −0.47 [−0.53, −0.45] | 0.90 [0.79, 1.06] | 0.04 [0.02, 0.06] | 1.48 [1.34, 1.58] | 1.42 [1.31, 1.52] |
| Δ native face slope P95 (°) | −1.82 [−2.94, −1.24] | −1.54 [−1.98, −1.06] | 1.10 [0.03, 2.01] | 0.02 [−0.05, 0.17] | 2.94 [2.73, 3.21] | 2.92 [2.61, 3.14] |
| Δ triangle density (mm⁻²) | −19.2 [−23.2, −17.4] | −20.0 [−22.3, −19.0] | 58.2 [54.3, 62.1] | −0.6 [−0.8, −0.3] | 78.8 [75.6, 81.6] | 79.3 [75.7, 82.0] |
| Δ median edge length (µm) | 299.8 [268.3, 341.9] | 339.6 [315.8, 359.0] | −121.1 [−140.5, −113.2] | 40.7 [18.7, 58.6] | −431.8 [−463.3, −410.5] | −472.0 [−486.2, −445.4] |
| Δ median triangle area (µm²) | 119519 [101123, 139315] | 130378 [118151, 139390] | −23108 [−26016, −19616] | 11223 [424, 21972] | −143758 [−159431, −131606] | −156119 [−163585, −144372] |

Values are medians [95% percentile bootstrap confidence intervals] from 1,000 specimen-level resamples. Column labels A → B indicate signed differences B minus A for the same specimen and stage. ΔE₀₀ and Wasserstein distances are unsigned; Wasserstein distances use the indicated channel units.

ΔE₀₀ compares scan-mean CIELAB vectors. Wasserstein distances compare separate area-weighted L*, a* and b* distributions. Geometric comparisons use scan-level summaries without assuming pointwise correspondence. Colours are uncalibrated, and denser sampling does not establish greater geometric accuracy.

Table B10. Matched between-scanner colour and geometric differences. Panel C.

Mature specimens · Baseline · n = 32 matched specimens per scanner pair

| Metric | iTERO → TRIOS3 | iTERO → TRIOS5 | iTERO → LABscanner | TRIOS3 → TRIOS5 | TRIOS3 → LABscanner | TRIOS5 → LABscanner |
| --- | --- | --- | --- | --- | --- | --- |
| Colour difference ΔE₀₀ | 6.17 [5.86, 6.55] | 5.13 [4.66, 5.59] | 15.34 [14.26, 17.00] | 4.59 [4.22, 4.88] | 10.27 [9.96, 11.60] | 12.35 [11.97, 13.85] |
| Δ mean L* | 0.23 [−0.73, 0.92] | 0.17 [−1.23, 1.50] | 7.85 [6.90, 9.55] | 0.22 [−0.46, 1.13] | 7.53 [6.76, 9.06] | 7.86 [6.20, 9.75] |
| Δ mean a* | −9.46 [−10.20, −7.67] | −14.38 [−14.97, −13.18] | −21.47 [−24.17, −19.41] | −4.49 [−5.69, −4.04] | −11.42 [−13.57, −10.52] | −7.31 [−8.44, −5.82] |
| Δ mean b* | 7.00 [6.72, 7.92] | −0.50 [−0.85, −0.14] | 13.02 [12.40, 13.46] | −7.89 [−8.37, −7.42] | 5.63 [4.97, 6.47] | 13.28 [12.89, 14.28] |
| Wasserstein L* | 2.22 [1.64, 2.75] | 1.97 [1.53, 2.53] | 7.85 [6.90, 9.56] | 1.50 [1.36, 2.16] | 7.53 [6.75, 9.07] | 7.86 [6.20, 9.74] |
| Wasserstein a* | 9.46 [7.69, 10.16] | 14.38 [13.18, 14.96] | 21.47 [19.34, 24.14] | 4.52 [4.06, 5.70] | 11.44 [10.55, 13.58] | 7.26 [5.78, 8.47] |
| Wasserstein b* | 6.99 [6.72, 7.92] | 0.91 [0.74, 1.09] | 13.05 [12.42, 13.51] | 7.89 [7.42, 8.40] | 5.66 [5.02, 6.49] | 13.29 [12.92, 14.31] |
| Δ RMS height 50 µm grid (µm) | −2.57 [−4.86, −2.22] | −2.51 [−4.38, −1.94] | −0.89 [−2.65, 0.27] | 0.18 [−0.30, 0.64] | 1.78 [1.08, 2.59] | 1.73 [1.15, 2.22] |
| Δ RMS height 100 µm grid (µm) | −2.44 [−4.81, −2.13] | −2.41 [−4.03, −1.89] | −0.93 [−2.89, 0.12] | 0.15 [−0.31, 0.64] | 1.51 [0.84, 2.36] | 1.43 [0.96, 1.99] |
| Δ RMS height 200 µm grid (µm) | −2.25 [−4.77, −1.88] | −2.38 [−3.87, −1.78] | −1.45 [−3.62, −0.35] | 0.07 [−0.32, 0.65] | 1.00 [0.21, 1.47] | 0.77 [0.38, 1.33] |
| Δ native mean absolute height (µm) | −1.88 [−3.66, −1.47] | −1.95 [−2.83, −1.43] | −0.28 [−2.03, 0.58] | 0.09 [−0.19, 0.51] | 1.50 [0.97, 2.08] | 1.51 [0.99, 1.92] |
| Δ native RMS height (µm) | −2.75 [−5.09, −2.28] | −2.68 [−4.46, −2.08] | −1.22 [−3.08, 0.17] | 0.19 [−0.31, 0.64] | 1.88 [1.05, 2.68] | 1.87 [1.23, 2.32] |
| Δ native height P95−P5 (µm) | −7.38 [−14.45, −5.71] | −7.11 [−12.20, −5.69] | −0.88 [−8.66, 2.11] | 0.49 [−0.60, 2.26] | 6.03 [3.78, 8.77] | 5.77 [3.77, 7.31] |
| Δ median native face slope (°) | −0.58 [−0.68, −0.50] | −0.43 [−0.48, −0.32] | 2.34 [2.24, 2.74] | 0.19 [0.15, 0.31] | 3.12 [2.88, 3.40] | 2.90 [2.75, 3.01] |
| Δ native face slope P95 (°) | −1.42 [−2.11, −1.16] | −0.66 [−1.15, −0.26] | 6.49 [5.59, 7.47] | 0.96 [0.47, 1.31] | 7.92 [7.17, 8.84] | 7.15 [6.50, 7.76] |
| Δ triangle density (mm⁻²) | −30.8 [−34.3, −28.9] | −31.5 [−32.7, −29.3] | 109.0 [103.1, 119.0] | 0.5 [0.1, 1.1] | 144.0 [137.0, 149.8] | 143.8 [136.9, 148.5] |
| Δ median edge length (µm) | 261.8 [247.5, 277.3] | 250.0 [236.5, 260.9] | −111.7 [−121.9, −101.0] | −15.5 [−29.7, −3.1] | −372.4 [−382.3, −365.0] | −355.6 [−363.7, −348.9] |
| Δ median triangle area (µm²) | 78916 [73074, 80642] | 71440 [66456, 74449] | −16309 [−17294, −14814] | −8915 [−13206, −1890] | −91999 [−96860, −90079] | −88170 [−90391, −82663] |

Values are medians [95% percentile bootstrap confidence intervals] from 1,000 specimen-level resamples. Column labels A → B indicate signed differences B minus A for the same specimen and stage. ΔE₀₀ and Wasserstein distances are unsigned; Wasserstein distances use the indicated channel units.

ΔE₀₀ compares scan-mean CIELAB vectors. Wasserstein distances compare separate area-weighted L*, a* and b* distributions. Geometric comparisons use scan-level summaries without assuming pointwise correspondence. Colours are uncalibrated, and denser sampling does not establish greater geometric accuracy.

Table B10. Matched between-scanner colour and geometric differences. Panel D.

Mature specimens · Partial cleaning · n = 32 matched specimens per scanner pair

| Metric | iTERO → TRIOS3 | iTERO → TRIOS5 | iTERO → LABscanner | TRIOS3 → TRIOS5 | TRIOS3 → LABscanner | TRIOS5 → LABscanner |
| --- | --- | --- | --- | --- | --- | --- |
| Colour difference ΔE₀₀ | 5.01 [4.84, 5.11] | 5.13 [4.96, 5.29] | 11.32 [11.03, 11.67] | 4.62 [4.52, 4.83] | 6.86 [6.61, 7.06] | 8.36 [8.02, 8.64] |
| Δ mean L* | −0.01 [−0.55, 0.89] | 1.07 [0.44, 1.54] | 1.48 [1.11, 3.17] | 1.29 [0.70, 1.53] | 1.62 [1.42, 2.23] | 0.81 [0.26, 1.53] |
| Δ mean a* | −4.45 [−5.17, −3.73] | −12.85 [−13.70, −12.37] | −13.99 [−15.01, −13.11] | −8.35 [−8.74, −7.96] | −9.63 [−10.43, −8.93] | −1.40 [−1.96, −0.20] |
| Δ mean b* | 6.71 [6.49, 7.04] | −0.69 [−1.04, −0.52] | 12.11 [11.22, 13.18] | −7.39 [−7.77, −7.07] | 5.43 [4.72, 6.03] | 12.79 [12.25, 13.22] |
| Wasserstein L* | 1.79 [1.49, 2.47] | 1.47 [1.21, 1.85] | 2.27 [1.53, 3.18] | 1.41 [1.12, 1.63] | 1.63 [1.43, 2.24] | 1.36 [0.86, 1.66] |
| Wasserstein a* | 4.46 [3.75, 5.17] | 12.85 [12.36, 13.67] | 14.01 [13.13, 15.06] | 8.34 [7.96, 8.75] | 9.65 [8.98, 10.44] | 1.72 [1.50, 2.47] |
| Wasserstein b* | 6.70 [6.47, 7.02] | 0.99 [0.73, 1.14] | 12.09 [11.25, 13.21] | 7.38 [7.11, 7.78] | 5.48 [4.74, 6.04] | 12.83 [12.33, 13.22] |
| Δ RMS height 50 µm grid (µm) | −2.23 [−4.26, −1.55] | −2.61 [−4.28, −1.94] | −0.55 [−2.20, 0.35] | −0.31 [−0.53, −0.10] | 1.72 [1.15, 2.63] | 2.11 [1.45, 2.65] |
| Δ RMS height 100 µm grid (µm) | −2.08 [−3.88, −1.48] | −2.35 [−4.10, −1.75] | −0.41 [−1.79, 0.46] | −0.32 [−0.50, −0.10] | 1.67 [1.07, 2.65] | 2.03 [1.38, 2.60] |
| Δ RMS height 200 µm grid (µm) | −1.70 [−3.72, −1.32] | −2.04 [−3.81, −1.70] | −0.50 [−1.52, 0.38] | −0.27 [−0.45, −0.03] | 1.53 [0.87, 2.44] | 1.82 [1.18, 2.46] |
| Δ native mean absolute height (µm) | −1.45 [−2.88, −1.12] | −1.60 [−3.11, −1.33] | −0.49 [−1.05, 0.44] | −0.19 [−0.40, −0.03] | 1.28 [0.84, 1.86] | 1.49 [1.05, 2.05] |
| Δ native RMS height (µm) | −2.56 [−4.53, −1.67] | −3.10 [−5.21, −2.06] | −0.79 [−2.97, −0.11] | −0.31 [−0.66, −0.10] | 1.76 [1.10, 2.61] | 2.15 [1.47, 2.67] |
| Δ native height P95−P5 (µm) | −5.72 [−12.29, −4.22] | −5.99 [−12.82, −5.64] | −0.32 [−3.98, 2.03] | −0.47 [−1.38, −0.02] | 6.30 [3.77, 8.22] | 6.48 [5.01, 8.25] |
| Δ median native face slope (°) | −0.51 [−0.58, −0.43] | −0.50 [−0.59, −0.44] | 1.08 [0.99, 1.26] | 0.02 [0.00, 0.04] | 1.61 [1.53, 1.72] | 1.59 [1.50, 1.69] |
| Δ native face slope P95 (°) | −1.14 [−1.82, −0.88] | −1.24 [−1.80, −1.02] | 2.55 [1.99, 2.72] | −0.09 [−0.20, −0.06] | 3.60 [3.45, 3.98] | 3.77 [3.58, 4.15] |
| Δ triangle density (mm⁻²) | −19.5 [−21.5, −17.5] | −20.4 [−22.5, −18.2] | 63.6 [61.2, 66.4] | −0.7 [−1.0, −0.5] | 83.7 [80.7, 88.1] | 84.6 [81.8, 88.9] |
| Δ median edge length (µm) | 294.0 [249.7, 314.4] | 324.7 [299.8, 343.0] | −129.4 [−145.1, −117.7] | 36.6 [27.1, 55.5] | −413.6 [−428.3, −400.1] | −449.7 [−467.2, −437.0] |
| Δ median triangle area (µm²) | 103661 [89105, 110073] | 119618 [111574, 128430] | −23758 [−26658, −21417] | 17538 [12622, 30066] | −126276 [−131902, −115789] | −143246 [−149473, −133951] |

Values are medians [95% percentile bootstrap confidence intervals] from 1,000 specimen-level resamples. Column labels A → B indicate signed differences B minus A for the same specimen and stage. ΔE₀₀ and Wasserstein distances are unsigned; Wasserstein distances use the indicated channel units.

ΔE₀₀ compares scan-mean CIELAB vectors. Wasserstein distances compare separate area-weighted L*, a* and b* distributions. Geometric comparisons use scan-level summaries without assuming pointwise correspondence. Colours are uncalibrated, and denser sampling does not establish greater geometric accuracy.

Table B10. Matched between-scanner colour and geometric differences. Panel E.

Mature specimens · Clean · n = 32 matched specimens per scanner pair

| Metric | iTERO → TRIOS3 | iTERO → TRIOS5 | iTERO → LABscanner | TRIOS3 → TRIOS5 | TRIOS3 → LABscanner | TRIOS5 → LABscanner |
| --- | --- | --- | --- | --- | --- | --- |
| Colour difference ΔE₀₀ | 4.84 [4.63, 5.07] | 5.71 [5.39, 6.05] | 11.39 [11.15, 12.00] | 4.92 [4.80, 5.14] | 7.32 [6.79, 7.61] | 8.69 [8.46, 9.16] |
| Δ mean L* | −0.17 [−1.75, 0.56] | 1.34 [−0.07, 1.97] | 1.42 [0.24, 3.00] | 1.90 [1.50, 2.06] | 1.83 [1.46, 2.56] | −0.22 [−0.54, 0.46] |
| Δ mean a* | −4.50 [−5.07, −4.02] | −13.02 [−13.84, −12.67] | −12.56 [−13.43, −11.56] | −8.81 [−9.24, −8.29] | −8.21 [−9.64, −7.83] | 0.42 [−0.36, 1.70] |
| Δ mean b* | 5.81 [5.51, 6.18] | −0.93 [−1.42, −0.67] | 12.26 [11.48, 12.75] | −6.78 [−7.48, −6.51] | 5.99 [5.19, 6.65] | 13.00 [12.34, 14.25] |
| Wasserstein L* | 2.04 [1.56, 2.44] | 1.94 [1.64, 2.59] | 2.19 [1.54, 2.99] | 1.89 [1.52, 2.04] | 1.84 [1.52, 2.55] | 1.10 [0.91, 2.01] |
| Wasserstein a* | 4.60 [4.20, 5.08] | 13.04 [12.73, 13.84] | 12.46 [11.56, 13.44] | 8.83 [8.30, 9.24] | 8.17 [7.80, 9.55] | 1.95 [1.44, 2.72] |
| Wasserstein b* | 5.81 [5.53, 6.21] | 1.15 [0.83, 1.50] | 12.28 [11.55, 12.84] | 6.82 [6.62, 7.52] | 6.07 [5.29, 6.80] | 13.14 [12.43, 14.43] |
| Δ RMS height 50 µm grid (µm) | −3.27 [−4.45, −1.92] | −3.28 [−5.56, −1.90] | −2.70 [−5.27, −1.02] | −0.25 [−0.52, 0.19] | 0.66 [−0.13, 0.96] | 0.89 [0.34, 1.15] |
| Δ RMS height 100 µm grid (µm) | −3.07 [−4.18, −1.90] | −3.10 [−5.21, −1.81] | −2.54 [−4.76, −1.02] | −0.25 [−0.53, 0.18] | 0.58 [−0.15, 0.89] | 0.84 [0.30, 1.07] |
| Δ RMS height 200 µm grid (µm) | −2.87 [−3.77, −1.89] | −2.87 [−3.92, −1.72] | −2.35 [−3.95, −1.10] | −0.28 [−0.52, 0.18] | 0.35 [−0.18, 0.77] | 0.66 [0.16, 0.85] |
| Δ native mean absolute height (µm) | −1.80 [−3.12, −1.25] | −1.84 [−3.08, −1.28] | −1.70 [−2.80, −0.50] | −0.06 [−0.44, 0.15] | 0.45 [−0.13, 0.76] | 0.74 [0.27, 0.91] |
| Δ native RMS height (µm) | −3.46 [−4.59, −2.05] | −3.46 [−6.45, −1.94] | −2.90 [−6.08, −1.09] | −0.31 [−0.52, 0.18] | 0.60 [−0.14, 1.00] | 0.89 [0.35, 1.20] |
| Δ native height P95−P5 (µm) | −8.17 [−11.80, −5.30] | −7.02 [−12.22, −4.70] | −5.23 [−10.00, −2.32] | −0.32 [−1.36, 0.62] | 2.52 [0.43, 3.89] | 2.81 [1.29, 3.74] |
| Δ median native face slope (°) | −0.50 [−0.53, −0.46] | −0.47 [−0.50, −0.42] | 1.00 [0.90, 1.06] | 0.03 [0.01, 0.05] | 1.51 [1.45, 1.58] | 1.49 [1.41, 1.55] |
| Δ native face slope P95 (°) | −1.41 [−1.83, −1.28] | −1.52 [−1.91, −1.26] | 1.40 [1.02, 1.77] | 0.02 [−0.11, 0.08] | 2.96 [2.82, 3.14] | 2.95 [2.81, 3.04] |
| Δ triangle density (mm⁻²) | −19.4 [−21.5, −17.1] | −20.6 [−21.7, −18.1] | 57.9 [56.3, 62.2] | −1.0 [−1.3, −0.5] | 78.3 [76.8, 79.3] | 78.9 [77.8, 79.5] |
| Δ median edge length (µm) | 287.6 [277.7, 304.0] | 344.0 [322.0, 355.5] | −122.9 [−139.9, −113.4] | 48.5 [34.7, 65.0] | −412.8 [−422.4, −397.4] | −472.7 [−479.6, −457.7] |
| Δ median triangle area (µm²) | 102712 [98097, 112235] | 129856 [120860, 136132] | −22979 [−26975, −20786] | 23432 [13811, 31557] | −128884 [−131249, −123984] | −152254 [−158444, −149143] |

Values are medians [95% percentile bootstrap confidence intervals] from 1,000 specimen-level resamples. Column labels A → B indicate signed differences B minus A for the same specimen and stage. ΔE₀₀ and Wasserstein distances are unsigned; Wasserstein distances use the indicated channel units.

ΔE₀₀ compares scan-mean CIELAB vectors. Wasserstein distances compare separate area-weighted L*, a* and b* distributions. Geometric comparisons use scan-level summaries without assuming pointwise correspondence. Colours are uncalibrated, and denser sampling does not establish greater geometric accuracy.